Multi Anode Photomultiplier Front End Read-Out (MAP-FER)

(1 blank line)

Codin GHEORGHE1, Dirk MEIER1, Philip PÅHLSSON1, Petter ØYA1, Arne FREDRIKSEN1

(2 blank lines)

1Integrated Detector Electronics AS (IDEAS), (Gjerdrumsvei 19, 0484 Oslo, Norway)

(2 blank lines)

The Multi-Anode Photomultiplier Front End Read-Out (MAP-FER) has been designed to meet the requirements of the SONDE detector assembly to be installed at the European Spallation Source. SONDE uses neutron sensitive scintillators and a large area array of Hamamatsu H8500c Multi-anode Photomultiplier tubes (MaPMT) for neutron detection. The MAP-FER allows one to acquire the pulse heights from from the MaPMT anodes and to detect the pulses that exceed the programmed threshold.

The MAP-FER contains four readout ICs of type IDE3465 [1]. Each IC acquires the charge from the MaPMT and delivers the trigger signal and the analogue peak value of the event.

MAP-FER detects charges above a pre-programmed threshold and sequentially clocks out analogue pulse height from each of the IDE3465 ICs. The read-out cycle for the analogue values can be as low as 10 µs per IC. The ICs on the MAP-FER can also be used to only detect a pulse on any of the MaPMT channels which is above a given threshold. The IDE3465 provides 8–bit DACs to set the threshold for each channel individually. Such the threshold of each channel can be adapted to the gain of the respective MaPMT channel. The digital outputs for each IDE3465 channel deliver a digital LVTTL (low voltage TTL) pulse for each detected event. The detection rate per individual channel is above 100 thousand events per second per channel. The MAP-FER is designed for a negative charge range from 0 to -100pC.

A temperature sensor is incorporated on the MAP-FER to monitor the module temperature over time.

The MAP-FER circuit board can be mounted behind the MaPMT and is slightly smaller than the MaPMT. The combined assembly allows tiling of many MaPMT in a 2-dimensional array as required for the SONDE detector.

(1 blank line)

[1] D. Meier, et al., 2013, “Development of an ASIC for charged particle counting with silicon radiation detectors”, Nuclear Science Symposium and Medical Imaging Conference (NSS/MIC), 2013 IEEE (2013), [doi: 10.1109/NSSMIC.2013.6829764]

(1 blank line)

E-mail of the corresponding author:

codin.gheorghe@ideas.no